ROI1
ROI2
ROI3
ROI4
Global Parameters
Energy scale (channel -> energy)
Offset
Linear scale
Quadratic scale (optional)
Resolution
Fano factor
Noise
N × Regions-of-interest (with M peaks)
Incomplete charge
Bremsstrahlung
Linear
Quadratic
Peak intensities
Peak 0
...
Peak M-1
Input parameters
Transition energies
Material composition
Coating layer
Thickness
Composition
Default detector parameters
Zero offset
Linear scale
Transmission function
Fano factor
Noise

---
Case 1:
   One or a few peaks spread over a small energy range.
Limitations:
Can't fit energy scale
Can't fit both fano factor and noise
Fit noise, assume nominal fano factor
Additional assumptions:
Energy scale from default detector
Fano factor from default detector

---
Case 2:
   Two or three distinct peaks over a broad range of energies
Definition:
A broad range of energies is defined as greater than 1/3 the full range of measured channels or 20.0/3.0 keV which ever is smaller.
Can't fit quadratic energy scale
Additional assumptions:
Quadratic term is from default detector (nominally 0.0)
Detector efficiency

---
Case 3:
   More than three distinct peaks well distributed over a broad range of energies
Definition:
A broad range of energies is defined as greater than 1/3 the full range of measured channels or 20.0/3.0 keV which ever is smaller.
None!
Assumptions:
Detector efficiency

---
BAM CRM:
The BAM CRM is thin film material consisting of
Case 4:
   Two spectra collected at 10 keV and 30 keV on the BAM CRM.
None!
Assumptions:
BAM certified efficiencies
Mass absorption coefficients

---
What are the challenges:
Can't model Bremsstrahlung unless we know the composition
Can't model Bremsstrahlung unless detector efficiency
Can't model
Bremsstrahlung model must bracket all peaks accurately

---
ROI
X-Ray Transition
Global
Energy Scale
Quadratic
0:	zero offset
1:	gain
2:	quadratic
Linear
0:	zero offset
1:	gain
Gain
0:	gain
None
Resolution mode
FanoNoise
0:	Fano factor
1:	Noise factor
FWHM At Mn Kα
0:	FWHM at Mn Kα
None
Family weights
Tabulated weights
0:	Weight element 1
1:	Weight element 2
...
N-1:	Weight element N
Incomplete Charge
0:	Amplitude
1:	Width
Bremsstrahlung
0:	Coefficient A
1:	Coefficient B
Transition Intensity
0:	Transition 1
1:	Transition 2
...
N-1:	Transition N